## New process flow
High aspect ratio process development project
step | Tool | Recipe | Notes | Purpose
1. Drop gauge | Drop gauge |  | Measure just outside C,N,E,S,W features
2. Ash | YES asher | Recipe 4 Descum | Ash immediately prior to deposition | Wafer clean
3. Oxide | PECVD2 | USER_SiO2_350_6u | 32 min for single wafer, 43 min for wafer + afterclean | Hard mask for DRIE etch
4. Drop gauge/Filmmetrics | Drop gauge |  | Measure just outside C,N,E,S,W features. Filmmetrics is a more accurate way to measure oxide thickness especially when oxide thickness is <5 um. | measure oxide thickness
5. Ash | YES asher | Recipe 4 Descum | Ash immediately prior to coating | Improve resist adhesion
6. Coat | Rite Track (coater side) | talk to Jean | AZ10XT, 9 um | HMDS, coat, bake to deposit photoresist
6.1 EBR | Rite Track (coater side) | talk to Jean |  | remove edge bead for ICP-F
6.2 Rehydrate |  | 30 min |  | Important!!! Don't skip this step otherwise photoresist will crack post develop
7. Drop gauge | Drop gauge |  | Measure just outside C,N,E,S,W features | measure resist thickness
8. Expose | ABM | HAR DRIE mask | 84s total, split into two 42 second intervals with at least 30s inbetween exposure. | Expose test pattern into resist
9. Develop | Rite Track (developer Track) | DEV6 |  | Dissolve exposed regions of resist in developer to transfer mask pattern into resist
10. Inspect | NIKON |  | Check for completed develop and flaws (NOTE: features with sizes less than 3um may not resolve.) | Visually confirm exposure and development of resist
11. Ash | YES asher | Recipe 4 Descum |  | cleans remnants of exposed-yet-undeveloped resist via a low-power oxygen plasma
12. Etch oxide hardmask | ICP-F | QT SiO2 Etch-C4F8/O2 20C 5mT, 10.5 min |  | Removal of oxide layer in exposed areas
13. Drop gauge | Drop gauge |  | Measure just outside C,N,E,S,W features | measure remaining resist thickness
14. Keyence | Keyence |  | Measure trench and hole depths of C,N,E,S,W features
14.1 SEM * | SEM | may not be necessary depending on if keyence works | Sacrifice 1 wafer, cut across features, Do cross-section SEM with 45 degree angle stage, coat gold (Au/Pd recipe 30 sec) to minimize charging effect | verify side wall profile of resist, SiO2 fully etched in trenches
15. SPTS-DRIE DOE etch | SPTS-DRIE | See SPTS-DRIE tab
16. Drop gauge | Drop gauge |  | Measure just outside C,N,E,S,W features | measure remaining mask thickness
17. SEM | SEM |  | Sacrifice 1 wafer, cut across features, Do cross-section SEM with 45 degree angle stage, coat gold (Au/Pd recipe 30 sec) to minimize charging effect | Check etch depth, side wall profile etc.
Note on B6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Ask Grace about drop gauge approach
Note on D8: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Need to ask Jean about 1 um feature size resolution. May need to do AZ1512 --> oxide -->AZ10XT stack. Using new mask might not be reasonable
Note on D12: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Talk to Jean, probably need do to exposure matrix
Note on C16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Talk to Grace, probably need 15 minute etch, 15 loops?
Note on D18: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Realistically 1 um trenches are probably too small but we’ll see

## batch 3 data
Wafer |  |  | 1. Drop gauge | [6 empty] | 4. Drop gauge | [6 empty] | 7. Drop gauge | [6 empty] | 10. Inspect | 13. Drop gauge | [6 empty] | 14. Keyence (5um depths) | [6 empty] | 16. Drop gauge | [6 empty] | 17. SEM | [5 empty] | Scaled trench depth |  |  | Etch Rate | Erosion | Selectivity | Aspect Ratio
| Completed | Next Step | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | Notes | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | Feature | Wafer thickness | 5um trench depth | 3um trench depth | 1um trench depth | Comments | 5 um | 3 um | 1 um | um/loop (Si) | um/loop (SiO2) | (SiO2/Si) | (W/D)
B01 | 4. Drop gauge | 5. Ash | 518.14 | 518.88 | 518.31 | 516.96 | 518.94 | 518.246 | 0.0019103 | 524.37 | 525.63 | 525.03 | 523.47 | 524.6 | 524.62 | 0.0020586
B02 | 4. Drop gauge | 5. Ash | 518.8 | 519.63 | 518.6 | 517.57 | 518.81 | 518.682 | 0.0019858 | 525.16 | 526.76 | 524.93 | 523.5 | 525.26 | 525.122 | 0.003104
B03 | 4. Drop gauge | 5. Ash | 519.47 | 520.14 | 519.07 | 517.83 | 519.8 | 519.262 | 0.0022243 | 525.4 | 526.78 | 525.25 | 523.94 | 526.42 | 525.558 | 0.0027019
B04 | 4. Drop gauge | 5. Ash | 519.34 | 521.44 | 519.22 | 517.61 | 518.87 | 519.296 | 0.0036877 | 525.66 | 527 | 526.18 | 524.14 | 525.98 | 525.792 | 0.0027197
B05 | 4. Drop gauge | 5. Ash | 518.09 | 518.04 | 518.07 | 515.21 | 517.11 | 517.304 | 0.0027837 | 523.5 | 523.68 | 523.18 | 522.83 | 523.86 | 523.41 | 0.00098393
B06 | 4. Drop gauge | 5. Ash | 517.34 | 517.49 | 517.25 | 516.45 | 517.74 | 517.254 | 0.001247 | 523.98 | 524.18 | 524.35 | 522.62 | 523.4 | 523.706 | 0.0016517
B07 | 4. Drop gauge | 5. Ash | 521.85 | 520.89 | 519.63 | 517.77 | 519.48 | 519.924 | 0.0039237 | 529.22 | 528.38 | 527.17 | 525.26 | 526.67 | 527.34 | 0.0037547
B08 | 3. Oxide | 4. Drop gauge | 519.02 | 519.86 | 518.36 | 517.5 | 519.45 | 518.838 | 0.0022743
B09 | 3. Oxide | 4. Drop gauge | 522.76 | 522.45 | 522.56 | 521.49 | 522.78 | 522.408 | 0.0012347
B10 | 3. Oxide | 4. Drop gauge | 523.01 | 522.4 | 522.3 | 521.7 | 522.75 | 522.432 | 0.0012538
B11 | 1. Drop gauge | 2. Ash | 522.69 | 522.45 | 522.74 | 522.25 | 522.79 | 522.584 | 0.00051666
B12 | 1. Drop gauge | 2. Ash | 524.18 | 524.24 | 524.5 | 522.75 | 523.72 | 523.878 | 0.0016702
Oxide dep:
[10 empty] | C | N | E | S | W | Mean | U
[9 empty] | B01 | 6.23 | 6.75 | 6.72 | 6.51 | 5.66 | 6.374 | 0.085504
[9 empty] | B02 | 6.36 | 7.13 | 6.33 | 5.93 | 6.45 | 6.44 | 0.093168
[9 empty] | B03 | 5.93 | 6.64 | 6.18 | 6.11 | 6.62 | 6.296 | 0.056385
[9 empty] | B04 | 6.32 | 5.56 | 6.96 | 6.53 | 7.11 | 6.496 | 0.1193
[9 empty] | B05 | 5.41 | 5.64 | 5.11 | 7.62 | 6.75 | 6.106 | 0.20554
[9 empty] | B06 | 6.64 | 6.69 | 7.1 | 6.17 | 5.66 | 6.452 | 0.11159
[9 empty] | B07 | 7.37 | 7.49 | 7.54 | 7.49 | 7.19 | 7.416 | 0.023598
B08
B09
B10
B11
B12
Oxide dep Filmetrics:
[10 empty] | Min | Max | Mean | Std. Dev. | Unif. [%]
[9 empty] | B01 | 5.872 | 6.092 | 5.941 | 0.04783 | 1.8
[9 empty] | B02 | 5.951 | 6.158 | 6.014 | 0.04488 | 1.7
[9 empty] | B03 | 5.922 | 6.147 | 6.001 | 0.04851 | 1.9
[9 empty] | B04 | 5.892 | 6.113 | 5.964 | 0.04743 | 1.9
[9 empty] | B05 | 5.939 | 6.164 | 6.02 | 0.04869 | 1.9
[9 empty] | B06 | 5.935 | 6.165 | 6.022 | 0.04971 | 1.9
[9 empty] | B07 | 5.894 | 6.106 | 5.956 | 0.0454 | 1.8
[9 empty] | B08 | 5.967 | 6.185 | 6.041 | 0.04728 | 1.8
[9 empty] | B09 | 5.976 | 6.187 | 6.042 | 0.04596 | 1.8
[9 empty] | B10 | 5.943 | 6.162 | 6.017 | 0.04754 | 1.8
B11
B12

## ICP-F
Run | wafer # | Date | Recipe | Etch time (min) | Batches/loops | DC bias jump time | He flow
1 | A2 |  | QT SiO2 Etch-C4F8/O2 | 15 | 2
2 | A3 |  | QT SiO2 Etch-C4F8/O2 | 12 | 2
3 | A4 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
4 | A5 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
5 | A6 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
6 | A7 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
7 | A8 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
8 | A9 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
9 | A10 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
10 | A11 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
11 | A12 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
12 | A13 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
13 | A14 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 2
14 | A15 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 3
15 | A16 |  | QT SiO2 Etch-C4F8/O2 | 10.5 | 3
16 | B01 |  |  | 15 | 15

## SPTS-DRIE
Run | wafer # | Date | Seq | Recipe | He flow at loop 10 | He flow at loop 300 | He flow at the end | Max reflect. power 2 at loop 10 | Max reflect. power 2 at loop 300 | Max reflect. power 2 at the end | Loop set 1: | Loop set 2: | Loop set 3:
1 | A5 | 2025-01-13 | MM_Deep_STD_-3seq | MM_Deep_Standard_-3 | 14.2 sccm | 14.5 sccm | 14.9 sccm |  |  |  | 500
2 | A6 | 2025-01-29 | MM_Deep_3x_-3seq | MM_Deep_3X_-3 | 13.6 sccm | 14.2 sccm |  |  |  |  | 262 | 483 | 200
3 | A7 | 2025-02-21 | MM_Deep_3x_-3seq | MM_Deep_3X_-3 | 17.0 sccm | 18.6 sccm |  |  |  |  | 262 | 165 | 0
4 | A7.2 | 2025-02-21 | MM_Deep_3x_-3seq | MM_Deep_3X_-3 | 15.6 sccm | 16.4 sccm | 17.0 sccm |  |  |  | 1 | 318 | 300
5 | A7b | 2025-02-26 | MM_Deep_3x_-3seq | MM_Deep_3X_-3 | 17.0 sccm |  | 17.5 sccm |  |  |  | 1 | 1 | 200
6 | A7c | 2025-02-26 | MM_Deep_3x_-3seq | MM_Deep_3X_-3 | 17.0 sccm |  | 16.9 sccm |  |  |  | 1 | 1 | 150
7 | A8 | 2025-02-28 | MM_Deep_3xwC_-3seq | MM_Deep_3XwC_-3 | 16.5 sccm | 16.7 sccm | 18.3 sccm | 185 W | 204 W | 198 W | 262 | 483 | 550
8 | A9 | 2025-03-04 | MD_DOE4_3x_-3seq | MD_DOE4_3x_-3 | 17.1 sccm |  | 17.1 sccm | 226 W |  | 209 W | 1 | 1 | 20
9 | A10 | 2025-03-04 | MD_DOE2_3x_-3seq | MD_DOE2_3x_-3 | 16.5 sccm | 16.8 sccm | 18.2 sccm | 196 W | 183 W | 206 W | 262 | 483 | 550
10 | A11 | 2025-03-05 | MD_DOE1_3x_-3seq | MD_DOE1_3x_-3 | 18.2 sccm | 19.5 sccm | 20.0 sccm | 194 W | 186 W | 174 W | 262 | 192 | 0
11 | A11b | 2025-03-05 | MD_DOE1_3x_-3seq | MD_DOE1_3x_-3 | 18.1 sccm | 19.3 sccm | 20.0 sccm | 189 W | 215 W | 239 W | 1 | 291 | 272
12 | A12 | 2025-03-06 | MD_DOE7_3x_-3seq | MD_DOE7_3x_-3 | 20.0 sccm | [5 empty] | 2 | 0 | 0
13 | A13 | 2025-03-06 | MD_DOE7_3x_-3seq | MD_DOE7_3x_-3 | 20.0 sccm | [5 empty] | 4 | 0 | 0
14 | A11c | 2025-03-08 | MD_DOE1_3x_-3seq | MD_DOE1_3x_-3 | 17.8 sccm |  | 17.9 sccm | 214 W |  | 223 W | 1 | 1 | 278
15 | A13 | 2025-03-08 | MD_DOE7_3x_-3seq | MD_DOE7_3x_-3 | 17.3 sccm | 17.7 sccm | 18.7 sccm | 200 W | 179 W | 239 W | 258 | 483 | 550
16 | A9 | 2025-03-10 | MD_DOE3_3x_-3seq | MD_DOE3_3x_-3 | 17.6 sccm | 18.1 sccm | 18.7 sccm | 183 W | 185 W | 177 W | 240 | 483 | 550
17 | A15 | 2025-03-11 | MD_DOE4_3x_-3seq | MD_DOE4_3x_-3 | 18.2 sccm | 18.8 sccm | 20.0 sccm | 192 W | 191 W | 176 W | 262 | 478 | 0
18 | A12 | 2025-03-13 | MD_DOE5_3x_-3seq | MD_DOE5_3x_-3 | 17.5 sccm | 17.8 sccm | 18.5 sccm | 193 W | 204 W | 178 W | 260 | 483 | 550
19 | A16 | 2025-03-14 | MD_DOE8_3x_-3seq | MD_DOE8_3x_-3 | 16.9 sccm | 17.5 sccm | 18.0 sccm | 201 W | 195 W | 175 W | 262 | 483 | 550
20 | A15 | 2025-03-16 | MD_DOE6_3x_-3seq | MD_DOE6_3x_-3 | 16.1 sccm | 16.8 sccm | 16.4 sccm | 181 W | 167 W | 173 W | 0 | 5 | 550
21 | A14 | 2025-03-16 | MD_DOE4_3x_-3seq | MD_DOE4_3x_-3 | 15.8 sccm | 16.6 sccm | 15.7 sccm | 196 W | 180 W | 240 W | 262 | 483 | 550

## 5 um depth
Drop gauge: |  | Keyence: | 500 | 200 | 100 | SEM50: | 500 | 200 | 100 | 50 | 25 | 15 | 10 | 5 | Drop gauge scaled: | 500 | 200 | 100 | 50 | 25 | 15 | 10 | 5 | Keyence scaled: | 500 | 200 | 100 | 50 | 25 | 15 | 10 | 5 | SEM200+ | 10 | 5 | Pseudo-depth: | 5
A9S | 1019.9 | A9S | 741.9 | 688.1 | 579.1 | 1087.5 | 753.8 | 716.3 | 615 | 513.8 | 373.1 | 153.8 | 121.9 | 131.3 | 1.0663 | 706.9431 | 671.7741 | 576.771 | 481.8617 | 349.9078 | 144.2397 | 114.3226 | 123.1383 | 1.0397 | 725.0366 | 688.9675 | 591.5329 | 494.1945 | 358.8633 | 147.9313 | 117.2486 | 126.2899 |  | 118.84 | 127.45 | A9S | 128.2383
A10S | 1012.7 | A10S | 817.7 | 724.9 | 635 | 1061.3 | 813.8 | 768.8 | 684.4 | 581.3 | 457.5 | 373.1 | 296.3 | 226.9 | 1.048 | 776.5337 | 733.5944 | 653.0593 | 554.6806 | 436.5498 | 356.0147 | 282.7316 | 216.5096 | 1.0445 | 779.1074 | 736.0258 | 655.2238 | 556.519 | 437.9966 | 357.1946 | 283.6686 | 217.2272 |  | 285.5 | 227.8 | A10S | 216.5096
A11S | 1019 | A11S | 831.2 | 746.3 | 656 | 1076.3 | 817.5 | 778.1 | 690 | 594.4 | 480 | 399.4 | 339.4 | 311.3 | 1.0562 | 773.978 | 736.6756 | 653.2658 | 562.7554 | 454.4458 | 378.1368 | 321.331 | 294.727 | 1.026 | 796.7947 | 758.3926 | 672.524 | 579.3453 | 467.8427 | 389.2842 | 330.8038 | 303.4155 |  | 333.8 | 304.2 | A11S | 294.727
A12C | 998.4 | A12C | 730.6 | 653.1 | 567.7 | 1061.3 | 740.6 | 705 | 615 | 521.3 | 414.4 | 341.3 | 281.3 | 251.3 | 1.063 | 696.7069 | 663.2168 | 578.5508 | 490.4041 | 389.8398 | 321.0722 | 264.6282 | 236.4062 | 1.0588 | 699.455 | 665.8328 | 580.8329 | 492.3385 | 391.3775 | 322.3386 | 265.672 | 237.3387 |  | 275.6 | 245.6 | A12C | 236.4062
A13S | 1017.1 | A13S | 816.5 | 732.3 | 638.3 | 1083.8 | 828.8 | 787.6 | 686.3 | 581.3 | 453.9 | 363.8 | 258.8 | 180 | 1.0656 | 777.7934 | 739.129 | 644.0632 | 545.5252 | 425.9658 | 341.4108 | 242.8727 | 168.9223 | 1.0553 | 785.399 | 746.3565 | 650.3611 | 550.8596 | 430.131 | 344.7492 | 245.2477 | 170.5741 |  | 253.6 | 174.8 | A13S | 168.9223
A14C | 929.82 | A14C | 697.1 | 610.3 | 515.7 | 969.4 | 720 | 631.9 | 543.8 | 444.4 | 311.3 | 221.3 | 120 | 82.5 | 1.0426 | 690.6028 | 606.0999 | 521.597 | 426.2554 | 298.5898 | 212.2645 | 115.1005 | 79.1316 | 1.0409 | 691.702 | 607.0646 | 522.4272 | 426.9339 | 299.0651 | 212.6023 | 115.2837 | 79.2575 |  | 115.8 | 79.2 | A14C | 176.2116
A15S | 971.38 | A15S | 704.6 | 633.3 | 551.2 | 1042.5 | 710.6 | 684.4 | 596.3 | 515.6 | 401.3 | 337.5 | 277.5 | 253.1 | 1.0732 | 662.1224 | 637.7098 | 555.62 | 480.4254 | 373.9231 | 314.4755 | 258.5688 | 235.8334 | 1.057 | 672.2747 | 647.4877 | 564.1393 | 487.7917 | 379.6564 | 319.2973 | 262.5334 | 239.4494 |  | 274.2 | 248.4 | A15S | 289.4534
A16C | 1022.5 | A16C | 738.4 | 662.4 | 578.6 | 1095 | 753.8 | 714.4 | 624.4 | 540 | 431.3 | 352.5 | 290.6 | 255 | 1.0709 | 703.8909 | 667.0995 | 583.0584 | 504.2466 | 402.7436 | 329.161 | 271.3594 | 238.1164 | 1.0595 | 711.4643 | 674.2772 | 589.3318 | 509.672 | 407.0769 | 332.7025 | 274.279 | 240.6784 |  | 281.3 | 244.7 | A16C | 238.1164

## DRIE etch rates
Wafer | Loops | Photoresist before: | Oxide before: | Total mask before: | Mask after: | Depth before: | Depth after: | Total etched: | Etch rate: | Erosion:
A5 | 500 | 2 | 3.975 | 5.975 | 2.2705 | 6.6818 | 299.362 | 296.3847 | 0.59277 | 0.0049241
A6 | 945 | 1.729 | 4.075 | 5.804 | 1.4384 | 6.2322 | 597.824 | 595.9574 | 0.63064 | 0.0032471
A7 | 1045 | 2.117 | 4.025 | 6.142 | 1.2174 | 6.5166 | 639.478 | 637.8859 | 0.61042 | 0.0031826
A7b | 200 | 0 | 1.2174 | 1.2174 | 0.47546 | 639.478 | 760 | 121.264 | 0.60632 | 0.0037099
A7c | 150 | 0 | 0.47546 | 0.47546 | 0.020536 | 760 | 850 | 90.4549 | 0.60303 | 0.0030328
A8 | 1295 | 2.3641 | 4.075 | 6.4391 | 0.47231 | 6.6924 | 800 | 799.2744 | 0.6172 | 0.0032365
DOE:
Wafer | Loops | Photoresist before: | Oxide before: | Total mask before: | Mask after: | Depth before: | Depth after: | Total etched: | Etch rate: | Erosion:
A9 | 1295 | 2.2486 | 4.025 | 6.2736 | 0.028995 | 6.6924 | 726.4667 | 726.0189 | 0.56063 | 0.0035614 | A11b | C | S | E
A10 | 1295 | 2.7292 | 4.075 | 6.8042 | 0.524 | 7.209 | 823.2333 | 822.3045 | 0.63498 | 0.0032726 | 500 | 654.8 | 652.9 | 658.9
A11 | 1295 | 1.6586 | 3.975 | 5.6336 | 0.1119 | 6.4888 | 834.4 | 833.4329 | 0.64358 | 0.0033091 | 200 | 588.8 | 591.3 | 592.9
A12 | 1295 | 1.8818 | 4.011 | 5.8928 |  | 6.3816 |  |  |  |  | 100 | 523.2 | 527.8 | 525.5
A13 | 1295 | 2.5596 | 4.091 | 6.6506 | 0.42772 | 7.1194 | 822.1333 | 821.2368 | 0.63416 | 0.003336
A14 | 1295 | 0 | 4.081 | 4.081 |  | 4.4538 |  |  |  |  | A11c | C | S | E
A15 | 1295 | 1.6949 | 4.091 | 5.7859 |  | 6.3468 |  |  |  |  | 500 | 829.8 | 831.2 | 842.2
A16 | 1295 | 2.6586 | 3.985 | 6.6436 |  | 7.1764 |  |  |  |  | 200 | 744 | 746.3 | 751.4
[11 empty] | 100 | 656.4 | 656 | 659.7

## Process flow
High aspect ratio process development project
step | Tool | Recipe | Notes | Purpose
0 Ash | YES asher | Recipe 4 Descum | Ash immediately prior to deposition | Wafer clean
1. Oxide | PECVD2 | Oxide: USER_SiO2_350_4u | Oxide: 426 nm/min, time = 9.4 min | Silicon oxide will be the hardmask for the DRIE etching
2 Filmetrics | Filmetrics | SiO2 on Si | Si reference wafer, edit expected thickness to 4000 nm | confirm SiO2 thickness
3. Ash | YES asher | Recipe 4 Descum | Ash immediately prior to coating | Improve resist adhesion
4 COAT | Rite Track (coater side) | 10XTD2 | AZ10XT, 7 um | HMDS, coat, bake to deposit photoresist
4.1 EBR | Rite Track (coater side) | TOPEBR20 |  | remove edge bead for ICP-F etcher (SiO2 etching)
5. REHYDRATE |  | 30 min |  | Important!!! Don't skip this step otherwise photoresist will crack post develop
6. FILMETRICS | Filmetrics | MFF AZ10xT on SiO2 | Si reference wafer, edit expected thickness to 7 um over 4 um | measure resist thickness
7. EXPOSE | ABM | 84s total, split into two 42 second intervals with at least 30s inbetween exposure. | Use WNF Uniformity Mask | Expose test pattern into resist
8. DEVELOP | Rite Track (developer Track) | DEV6 |  | Dissolve exposed regions of resist in developer to transfer mask pattern into resist
9. INSPECT | NIKON |  | Check for completed develop and flaws (NOTE: features with sizes less than 3um may not resolve.) | Visually confirm exposure and development of resist
10. DESCUM | YES asher | Recipe 4 Descum | Recipe 4 | cleans remnants of exposed-yet-undeveloped resist via a low-power oxygen plasma
11. DEKTAK | Dektak | Standard scan, range = 65.5 um, hills and valleys, length = 500 um, resolution = 0.1 | Measure 500 um trench 270 degrees for S,E,C (intuitive coordinates) 90 degrees for N and W (inverse coordinates) | Measure resist thickness for direct comparison to post etching measurement. NOTE: This tool has a minimum trench width capability and typically is used to measure only the 500um or 200um wide trench depth. | measure south, east, center at 270 degrees (intuitive coordinates), then flip to 90 degrees and measure north and west (inverse)
11.1 SEM* | SEM |  | Sacrifice 1 wafer, cut across features, Do cross-section SEM with 45 degree angle stage, coat gold (Au/Pd recipe 30 sec) to minimize charging effect | verify side wall profile of resist --> looks good with current setup✅
12. Etch oxide hardmask | ICP-F | QT SiO2 Etch-C4F8/O2 20C 5mT, 10.5 min | To avoid reticulation, rest for one minute in chamber after 5m15s | Removal of the SiO2 layer in the exposed areas
13. Nanospec | Nanospec | WNF folder: Grace -AZ10xT onSiO2 on Si | Edit recipe --> 2 um on 4 um, measure unetched area in five different spots | Measure remaining photoresist thickness
14. Dektak | Dektak | Standard scan, range = 65.5 um, hills and valleys, length = 500 um, resolution = 0.1 | Measure 500 um trench 270 degrees for S,E,C (intuitive coordinates) 90 degrees for W (inverse coordinates), 180 degrees for major flat |  | due to ICP clamp, should rotate 0 degrees for major flat measurement and find clean feature
14.1 SEM * | SEM |  | Sacrifice 1 wafer, cut across features, Do cross-section SEM with 45 degree angle stage, coat gold (Au/Pd recipe 30 sec) to minimize charging effect | verify side wall profile of resist, SiO2 fully etched in trenches --> looks good with current setup✅
15. SPTS-DRIE DOE etch | SPTS-DRIE | See SPTS-DRIE tab | Target: 500 um
16. Nanospec | Nanospec | MFF SiO2 on Si | Measure unetched area | SiO2 thickness
17. Dektak | Dektak | Standard scan, range = 1 mm, valleys, length = 600 um, resolution = 0.1 | Measure 500 um trench | SiO2 + trench depth
18. SEM | SEM |  | Sacrifice 1 wafer, cut across features, Do cross-section SEM with 45 degree angle stage, coat gold (Au/Pd recipe 30 sec) to minimize charging effect | Verify etch depth, check side wall profile etc.
For second batch:
17. Keyence | Keyence |  | Measure 500 um, 200 um, and 100 um trench of C, S, E features | SiO2 + trench depth

## Drop gauge
| C | E | S | Mean | U
Assumption before: | 1025 | 1025 | 1025 | 1025 | 0
A14 before: | 1026.9 | 1022.7 | 1019.1 | 1022.9 | 0.0038127
A9 after: | 1001.1 | 1006.6 | 1019.9 | 1009.2 | 0.0093143
A10 after: | 1020.2 | 1040.6 | 1012.7 | 1024.5 | 0.013616
A11 after: | 1034.5 | 1035.08 | 1019 | 1029.5267 | 0.0078094
A12 after: | 998.4 | 1024.4 | 1004.1 | 1008.9667 | 0.012884
A13 after: | 1022.1 | 1025.2 | 1017.1 | 1021.4667 | 0.0039649
A14 after: | 929.82 | 921.98 | 918.9 | 923.5667 | 0.0059119
A15 after: | 963.8 | 975.63 | 971.38 | 970.27 | 0.0060962
A16 after: | 1022.5 | 985.8 | 925.7 | 978 | 0.049489
A9 etched: | 23.9 | 18.4 | 5.1 | 15.8 | 0.59494
A10 etched: | 4.8 | -15.6 | 12.3 | 0.5 | 27.9
A11 etched: | -9.5 | -10.08 | 6 | -4.5267 | -1.7761
A12 etched: | 26.6 | 0.6 | 20.9 | 16.0333 | 0.81081
A13 etched: | 2.9 | -0.2 | 7.9 | 3.5333 | 1.1462
A14 etched: | 97.08 | 100.72 | 100.2 | 99.3333 | 0.018322
A15 etched: | 61.2 | 49.37 | 53.62 | 54.73 | 0.10808
A16 etched: | 2.5 | 39.2 | 99.3 | 47 | 1.0298

## batch 1 data
Wafer |  |  | 2 Filmetrics |  |  |  |  | 6 Filmetrics |  |  |  |  | 11 DEKTAK | [6 empty] | 12 ETCH |  |  |  | 13 NANOSPEC | [7 empty] | 14 DEKTAK | [6 empty] | 15 ETCH |  |  |  |  | 16 NANOSPEC | [7 empty] | 17 DEKTAK | [6 empty] | 18 SEM |  |  |  |  | Etch Rate | Erosion | Selectivity | Aspect Ratio
| Completed | Next Step | Min | Max | Mean | Std. Dev. | Unif. [%] | Min | Max | Mean | Std. Dev. | Unif. [%] | C | N (major flat) | E | S | W | Mean | U | Tool/Operator | Recipe | Etch time | Date | Recipe | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | Tool/Operator | Recipe | He (sccm@torr) or bias/CL/CT/APC | Loops or min | Date | Recipe | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | Step: | 500um depth | 5um depth | 5um width | Comments | um/loop | um/loop (SiO2) | (SiO2/Si) | (in 5um trench)
A1 | 11.1 SEM* |  | 4000 | 4050 | 4025 |  |  | 7 | 7.3 | 7.15 |  |  | 7.071 | 7.214 | 7.183 | 7.194 | 7.208 | 7.174 | 0.0099665 | [39 empty] | 11.1 SEM* |  |  |  | everything looks good
A2 | 14.1 SEM* |  | 4050 | 4100 | 4075 |  |  | 7 | 7.4 | 7.2 |  |  | 7.126 | 7.389 | 7.249 | 7.253 | 7.213 | 7.246 | 0.018148 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*7.5 min | 2024-12-20 | [8 empty] | 5.495 | 5.834 | 5.565 | 5.676 | 5.8 | 5.674 | 0.029873 | [20 empty] | 14.1 SEM* |  |  |  | probably overetched
A3 | 14.1 SEM* |  | 4050 | 4100 | 4075 |  |  | 7.1 | 7.4 | 7.25 |  |  | 7.053 | 7.199 | 7.056 | 7.184 | 7.223 | 7.143 | 0.0119 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*6 min | 2024-12-20 | [8 empty] | 6.526 | 6.521 | 6.511 | 6.595 | 6.528 | 6.5362 | 0.0064258 | [20 empty] | 14.1 SEM* |  |  |  | 20% overetched in 5um feature
A4 | 14.1 SEM* |  | 4050 | 4100 | 4075 |  |  | 7 | 7.3 | 7.15 |  |  | 6.999 | 7.245 | 7.141 | 7.208 | 7.194 | 7.1574 | 0.017185 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-02 | [35 empty] | 14.1 SEM* |  |  |  | 5-15% overetched
A5 | 18. SEM |  | 3950 | 4000 | 3975 |  |  | 6.9 | 7.2 | 7.05 |  |  | 6.997 | 7.206 | 7.195 | 7.255 | 7.335 | 7.1976 | 0.02348 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-02 | [6 empty] | 2 |  | 6.582 | 6.721 | 6.711 | 6.704 | 6.691 | 6.6818 | 0.010401 | SPTS-DRIE/MD | MM_Deep_Standard_-3 | 14.5 sccm | 500 | 2025-01-13 | MFF SiO2 on Si | 2270.5 |  |  |  |  | 2270.5 | 0 | 300.12 | 297.62 | 301.43 | 298.41 | 299.23 | 299.362 | 0.0063635 | 18. SEM | 303.8 | 131.4 | 5 | pillars toppled, difficulty etching smaller trenches | 0.59418 | 0.0049241 | 120.6681 | 25.8847
A6 | 18. SEM |  | 4050 | 4100 | 4075 |  |  | 6.9 | 7.5 | 7.2 |  |  | 7.066 | 7.165 | 7.146 | 7.172 | 7.183 | 7.1464 | 0.0081859 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-20 | Grace -AZ10xT onSiO2 on Si | 1.5332 | 1.7725 | 1.6936 | 1.853 | 1.7928 | 1.729 | 0.09248 | 6.166 | 6.152 | 6.205 | 6.39 | 6.248 | 6.2322 | 0.019094 | SPTS-DRIE/MD | MM_Deep_3X_-3 | 14.2 sccm | 945 | 2025-01-29 | MFF SiO2 on Si | 1273.1 | 1609.7 | 1394.4 | 1339.9 | 1574.8 | 1438.38 | 0.11701 | 600.03 | 595 | 600.27 | 594.75 | 599.07 | 597.824 | 0.0046167 | 18. SEM | 613.1 | 234.6 | 5 | 500 um trench has angled sidewalls | 0.63064 | 0.0032471 | 194.2187 | 45.8463
A7 |  |  | 4000 | 4050 | 4025 |  |  | 7.1 | 7.4 | 7.25 |  |  | 7.044 | 7.185 | 7.109 | 7.075 | 7.115 | 7.1056 | 0.0099218 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-20 | Grace -AZ10xT onSiO2 on Si | 1.8881 | 2.1422 | 2.1121 | 2.2328 | 2.2096 | 2.117 | 0.081414 | 6.44 | 6.415 | 6.532 | 6.638 | 6.558 | 6.5166 | 0.01711 | SPTS-DRIE/MD | MM_Deep_3X_-3 | 17.0 sccm | 1045 | 2025-02-21 | MFF SiO2 on Si | 1144.4 | 1152.1 | 1216.1 | 1314.5 | 1260.1 | 1217.44 | 0.06986 | 642.49 | 637.62 | 645.63 | 632.55 | 639.1 | 639.478 | 0.010227 | [5 empty] | 0.61078 | 0.0031826 | 191.9104
A7b | [38 empty] | SPTS-DRIE/MD | MM_Deep_3X_-3 | 17.0 sccm | 200 | 2025-02-26 | MFF SiO2 on Si | 388.89 | 404.65 | 494.43 | 594.53 | 494.8 | 475.46 | 0.21625 | From Keyence, depth was between 750-800 um |  |  |  |  | 760
A7c | 18. SEM | [37 empty] | SPTS-DRIE/MD | MM_Deep_3X_-3 | 17.0 sccm | 150 | 2025-02-26 | MFF SiO2 on Si | 0 | 0 | 0 | 51.01 | 51.67 | 20.536 | 1.258 | From Keyence, depth was about 850 um |  |  |  |  | 850 |  | 18. SEM |  |  |  | SWW feature
A8 | 18. SEM |  | 4050 | 4100 | 4075 |  |  | 7.1 | 7.6 | 7.35 |  |  | 7.108 | 7.171 | 7.148 | 7.175 | 7.145 | 7.1494 | 0.0046857 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-20 | Grace -AZ10xT onSiO2 on Si | 2.1839 | 2.2987 | 2.4897 | 2.4646 | 2.3837 | 2.3641 | 0.064675 | 6.674 | 6.607 | 6.664 | 6.847 | 6.67 | 6.6924 | 0.017931 | SPTS-DRIE/MD | MM_Deep_3XwC_-3 | 16.5 sccm | 1295 | 2025-02-28 | MFF SiO2 on Si | 383.81 | 369.38 | 487.03 | 609.43 | 511.88 | 472.306 | 0.25413 | From Keyence, depth was about 800 um |  |  |  |  | 800 |  | 18. SEM

## DOE data
Wafer |  |  | 2 Filmetrics |  |  |  |  | 6 Filmetrics |  |  |  |  | 11 DEKTAK | [6 empty] | 12 ETCH |  |  |  | 13 NANOSPEC | [7 empty] | 14 DEKTAK | [6 empty] | 15 ETCH |  |  |  |  | 16 NANOSPEC | [7 empty] | 17 Keyence (500 um) |  |  |  |  | 17 Keyence (200 um) |  |  |  |  | 17 Keyence (100 um) |  |  |  |  | 18 SEM |  |  |  |  | Etch Rate | Erosion | Selectivity | Aspect Ratio
| Completed | Next Step | Min | Max | Mean | Std. Dev. | Unif. [%] | Min | Max | Mean | Std. Dev. | Unif. [%] | C | N (major flat) | E | S | W | Mean | U | Tool/Operator | Recipe | Etch time | Date | Recipe | C | N (major flat) | E | S | W | Mean | U | C | N (major flat) | E | S | W | Mean | U | Tool/Operator | Recipe | He (sccm@torr) or bias/CL/CT/APC | Loops or min | Date | Recipe | C | N (major flat) | E | S | W | Mean | U | C | S | E | Mean | U | C | S | E | Mean | U | C | S | E | Mean | U | 500um depth | 200um depth | 100um depth | 5um depth (200+) | Comments | um/loop | um/loop (SiO2) | (SiO2/Si) | (in 5um trench)
A7 | 18. SEM |  | 4000 | 4050 | 4025 |  |  | 7.1 | 7.4 | 7.25 |  |  | 7.044 | 7.185 | 7.109 | 7.075 | 7.115 | 7.1056 | 0.0099218 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-20 | Grace -AZ10xT onSiO2 on Si | 1.8881 | 2.1422 | 2.1121 | 2.2328 | 2.2096 | 2.117 | 0.081414 | 6.44 | 6.415 | 6.532 | 6.638 | 6.558 | 6.5166 | 0.01711 | SPTS-DRIE/MD | MM_Deep_3X_-3 | 17.0 sccm | 150 | 2025-02-26 | MFF SiO2 on Si | 0 | 0 | 0 | 51.01 | 51.67 | 20.536 | 1.258 |  | 845 |  | 845 | 0 |  | 746 |  | 746 | 0 |  | 655 |  | 655 | 0 | 860.8 | 772.2 | 683.5 | 304.9 | SWW feature
A8 | 18. SEM |  | 4050 | 4100 | 4075 |  |  | 7.1 | 7.6 | 7.35 |  |  | 7.108 | 7.171 | 7.148 | 7.175 | 7.145 | 7.1494 | 0.0046857 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-01-20 | Grace -AZ10xT onSiO2 on Si | 2.1839 | 2.2987 | 2.4897 | 2.4646 | 2.3837 | 2.3641 | 0.064675 | 6.674 | 6.607 | 6.664 | 6.847 | 6.67 | 6.6924 | 0.017931 | SPTS-DRIE/MD | MM_Deep_3XwC_-3 | 16.5 sccm | 1295 | 2025-02-28 | MFF SiO2 on Si | 383.81 | 369.38 | 487.03 | 609.43 | 511.88 | 472.306 | 0.25413 | 804 | 797.5 | 804.2 | 801.9 | 0.0041776 | 711.3 | 708.6 | 713.1 | 711 | 0.0041776 | 621.6 | 624.4 | 625 | 623.6667 | 0.0031646 | 820.3 | 744.9 | 661.5 | 292.5 | C feature
A9 | 18. SEM |  | 3928 | 4074 | 3975 | 44.17 | 1.8 | 6.31 | 6.769 | 6.598 | 0.08204 | 3.5 | 6.71 | 6.59 | 6.65 | 6.96 | 6.6 | 6.702 | 0.027604 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-02-03 | Grace -AZ10xT onSiO2 on Si | [7 empty] | 6.879 | 6.559 | 6.661 | 6.732 | 6.631 | 6.6924 | 0.023908 | SPTS-DRIE/MD | MD_DOE3_3x_-3 | 17.1 sccm | 1295 | 2025-03-10 | MFF SiO2 on Si | 37.26 | 25.55 | 26.75 |  | 26.42 | 28.995 | 0.20193 | 713.8 | 741.9 | 723.7 | 726.4667 | 0.01934 | 637.3 | 688.1 | 646.2 | 657.2 | 0.01934 | 551.5 | 579.1 | 561.2 | 563.9333 | 0.038649 | 806.3 |  |  |  | S feature
A10 | 18. SEM |  | 3969 | 4102 | 4011 | 40.04 | 1.7 | 6.906 | 7.335 | 7.171 | 0.0998 | 3 | 7.08 | 7.2 | 7.19 | 7.22 | 7.21 | 7.18 | 0.0097493 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-02-03 | Grace -AZ10xT onSiO2 on Si | [7 empty] | 6.724 | 7.666 | 6.895 | 7.276 | 7.484 | 7.209 | 0.065335 | SPTS-DRIE/MD | MD_DOE2_3x_-3 | 16.5 sccm | 1295 | 2025-03-04 | MFF SiO2 on Si |  |  | 524 |  |  | 524 |  | 824.2 | 817.7 | 827.8 | 823.2333 | 0.0061343 | 727.6 | 724.9 | 730.3 | 727.6 | 0.0061343 | 636.5 | 635 | 639.8 | 637.1 | 0.0037108 | 837.5 | 762.6 | 680 | 221.5 | CW feature
A11 | 18. SEM |  | 4033 | 4236 | 4091 | 62.74 | 2.5 | 6.847 | 7.451 | 7.232 | 0.1288 | 4.2 | 7.021 | 7.204 | 7.177 | 7.203 | 7.234 | 7.1678 | 0.014858 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-02-10 | Grace -AZ10xT onSiO2 on Si | 1.5273 |  | 1.7365 | 1.712 |  | 1.6586 | 0.063065 | 6.156 | 6.367 | 6.266 | 6.271 | 7.384 | 6.4888 | 0.094625 | SPTS-DRIE/MD | MD_DOE1_3x_-3 | 17.8 sccm | 1295 | 2025-03-08 | MFF SiO2 on Si | 9.23 |  | 160.65 | 165.82 |  | 111.9 | 0.69969 | 829.8 | 831.2 | 842.2 | 834.4 | 0.0074305 | 744 | 746.3 | 751.4 | 747.2333 | 0.0074305 | 656.4 | 656 | 659.7 | 657.3667 | 0.0049516 |  |  |  |  | S feature
A12 | 18. SEM |  | 4034 | 4212 | 4081 | 50.49 | 2.2 | 6.893 | 7.42 | 7.222 | 0.1156 | 3.6 | 6.986 | 7.213 | 7.135 | 7.208 | 7.234 | 7.1552 | 0.01733 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-02-11 | Grace -AZ10xT onSiO2 on Si | 1.7451 | 1.8776 | 1.9382 | 1.9105 | 1.9377 | 1.8818 | 0.051307 | 6.305 | 6.401 | 6.449 | 6.343 | 6.41 | 6.3816 | 0.011282 | SPTS-DRIE/MD | MD_DOE5_3x_-3 | 17.5 sccm | 1295 | 2025-03-13 | MFF SiO2 on Si | 0 | 15.26 | 15.21 | 16.8 | 15.46 | 12.546 | 0.66954 | 730.6 | 757.2 | 760 | 749.2667 | 0.019619 | 653.1 | 682.5 | 680.5 | 672.0333 | 0.019619 | 567.7 | 599.6 | 594.7 | 587.3333 | 0.021874
A13 | 18. SEM |  | 4034 | 4247 | 4091 | 64.96 | 2.6 | 7.073 | 7.464 | 7.258 | 0.1023 | 2.7 | 7.056 | 7.223 | 7.194 | 7.186 | 7.292 | 7.1902 | 0.016411 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-02-11 | Grace -AZ10xT onSiO2 on Si | [7 empty] | 7.05 | 7.217 | 7.058 | 7.103 | 7.169 | 7.1194 | 0.011729 | SPTS-DRIE/MD | MD_DOE7_3x_-3 | 17.3 sccm | 1295 | 2025-03-06 | MFF SiO2 on Si | 321.3 |  | 476.24 | 485.62 |  | 427.72 | 0.19209 | 823.5 | 816.5 | 826.4 | 822.1333 | 0.0060209 | 733.6 | 732.3 | 737.7 | 734.5333 | 0.0060209 | 640.4 | 638.3 | 644.7 | 641.1333 | 0.0036758
A14 | 18. SEM |  | 3938 | 4090 | 3985 | 41.17 | 1.9 | 6.814 | 7.397 | 7.17 | 0.1542 | 4.1 | 7.124 | 7.243 | 7.239 | 7.313 | 7.271 | 7.238 | 0.013056 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 2*5.25 min | 2025-02-11 | Grace -AZ10xT onSiO2 on Si | [7 empty] | 4.605 | 4.429 | 4.489 | 4.403 | 4.343 | 4.4538 | 0.029413 | SPTS-DRIE/MD | MD_DOE4_3x_-3 | 15.8 sccm | 1295 | 2025-03-16 | MFF SiO2 on Si | 0 | 110.09 | 0 | 0 | 0 | 22.018 | 2.5 | 697.1 | 696.7 | 706 | 699.9333 | 0.0066435 | 610.3 | 615.9 | 621.1 | 615.7667 | 0.0066435 | 515.7 | 520.8 | 529.3 | 521.9333 | 0.0087696
A15 | 18. SEM |  | 3926 | 4039 | 3953 | 22.61 | 1.4 | 7.022 | 7.252 | 7.186 | 0.06452 | 1.6 | 7.317 | 7.254 | 7.571 | 7.442 | 7.322 | 7.3812 | 0.021473 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 3*3.5 min | 2025-02-16 | Grace -AZ10xT onSiO2 on Si | 1.5678 | 1.7647 | 1.7057 | 1.7416 |  | 1.6949 | 0.058084 | 6.204 | 6.238 | 6.349 | 6.38 | 6.563 | 6.3468 | 0.028282 | SPTS-DRIE/MD | MD_DOE6_3x_-3 | 16.1 sccm | 1295 | 2025-03-16 | MFF SiO2 on Si | 0 | 0 | 0 | 0 | 27.7 | 5.54 | 2.5 | 679.5 | 704.6 | 704.9 | 696.3333 | 0.018238 | 603 | 633.3 | 626.1 | 620.8 | 0.018238 | 518.7 | 551.2 | 540.8 | 536.9 | 0.024404
A16 | 18. SEM |  | 4025 | 4126 | 4049 | 19.75 | 1.2 | 7.067 | 7.597 | 7.255 | 0.1258 | 3.6 | 7.331 | 7.553 | 7.596 | 7.346 | 7.357 | 7.4366 | 0.017817 | ICP-F/MD | QT SiO2 Etch-C4F8/O2 | 3*3.5 min | 2025-02-16 | Grace -AZ10xT onSiO2 on Si | [7 empty] | 6.951 | 7.454 | 7.15 | 7.2 | 7.127 | 7.1764 | 0.035045 | SPTS-DRIE/MD | MD_DOE8_3x_-3 | 16.9 sccm | 1295 | 2025-03-14 | MFF SiO2 on Si | 132.89 |  | 50.44 |  |  | 91.665 | 0.44974 | 738.4 | 625.3 | 737.6 | 700.4333 | 0.080736 | 662.4 | 561.7 | 661.5 | 628.5333 | 0.080736 | 578.6 | 484.3 | 579.1 | 547.3333 | 0.080107
8.3992

## DOE
Current variations from loop set to loop set:
| Set 1 | Set 2 | Set 3
Dep time | 1.2 | 1.2-->1.4 | 1.4
Etch 1 time | 1.2 | 1.2-->1.5 | 1.5
Etch 1 HF Power | 64-->70 | 70-->80 | 80-->60
Etch 2 HF Power | 24 | 24 | 20
Dep HF Tune Capaciter | 20 | 25 | 20
Etch 1 HF Tune Capaciter | 20 | 25 | 20
Etch 2 HF Tune Capaciter | 25 | 20 | 20
Dep HF Load Capaciter | 30 | 25 | 30
Etch 1 HF Load Capaciter | 30 | 25 | 30
Etch 2 HF Load Capaciter | 25 | 30 | 30

## MM_Deep_3x_-3
Recipe name |  | MM_Deep_3x_-3
| Unit | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Step name |  | Strike 1 | Dep 1 | Etch 1_1 | Etch 1_2 | Dep 2 | Etch 2_1 | Etch 2_2 | Dep 3 | Etch 3_1 | Etch 3_2
Process Time | secs | 1 | 1.2 | 1.2 | 2.5 2.5 | 1.2 1.4 | 1.2 1.5 | 2.5 2.5 | 1.4 1.4 | 1.5 1.5 | 2.5 2.5
Process Pressure | mTorr | 40.0 +/- 0% | 40.0 +/- 50% | 30.0 30.0 +/- 50% | 30.0 30.0 +/- 50% | 40.0 40 +/- 50% | 30.0 30 +/- 50% | 30.0 30.0 +/- 50% | 40.0 40.0 +/- 50% | 30.0 30.0 +/- 50% | 30.0 30.0 +/-50%
APC Setpoint Position | % | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
APC Mode |  | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic
Source power | Watts | 2500.0 +/- 0 | 2500.0 +/- 50% | 2500.0 +/- 50% | 2500.0 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50%
Source MU Tune capaciter |  | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
Source MU load capaciter |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source RF Control Mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Source 2 Power | Watts | 500.0 +/- 0 | 500.0 +/- 50% | 500.0 +/- 50% | 500.0 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50%
Source 2 MU tune capaciter |  | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source 2 MU load capaciter |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source 2 RF Control mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Platen HF Power | Watts | 0.0 +/- 0% | 0.0 +/- 0% | 64.0 70.0 +/- 30% | 24 +/- 40% | 0.0 0.0 +/- 0% | 70.0 80 +/- 30% | 24.0 24 +/- 40% | 0.0 0.0 +/- 0% | 80.0 60.0 +/-30% | 20.0 20.0 +/-40$
Platen HF Tune Capaciter Adjust |  | Present | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic
Platen HF Tune Capacter |  | 30.0 +/- 20% | 20.0 +/- 5% | 20.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 20.0 +/- 5% | 20.0 +/- 5% | 20.0 +/-5% | 20.0 +/-5%
Platen HF load Capaciter |  | 60.0 +/- 20% | 30.0 +/- 5% | 30.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5%
Platen HF Padding Capaciter |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Platen HF control mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Platen HF modulation Enabled |  | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable
Platen HF Modulation Frequence | Hz | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150
Platen HF Modulation Duty Cycle | % | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Hellium Pressure | Torr | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/-20 %
Hellium Flow warning level | sccm | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Hellium Flow fault level | sccm | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25
Coil current | Amps | 10.0 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5%
Loop destination |  | 0 | 0 | 0 | 2 | 0 | 0 | 5 | 0 | 0 | 8
Number of loops |  | 0 | 0 | 0 | 262 | 0 | 0 | 483 | 0 | 0 | 550
Loop Variation Parameter |  | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Gas line Config |  | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow
P1 Argon | sccm | 300 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P2 Oxygen | sccm | 100 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P3 Nitrogen | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P4 C4F8 | sccm | 220.0 +/- 20% | 220.0 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 220.0 220 +/-20% | 1.0 1.0 +/-0% | 1.0 1.0 +/- 0% | 220.0 220 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0%
P5 SF6 | sccm | 1.0 +/- 0% | 1.0 +/- 0% | 400.0 400 +/- 20% | 400 +/- 20% | 1.0 1.0 +/-0% | 400.0 400.0+/- 20% | 400.0 400.0 +/- 20% | 1.0 1.0 +/- 0% | 400.0 400.0 +/- 20% | 400.0 400.0 +/- 20%
S1 Oxygen | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
S2 C4F8 | sccm | 110.0 +/- 20% | 110.0 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 110.0 110 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 110.0 110 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0%
S3 SF6 | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%

## Important Resources
Batch 1 powerpoint link:
DRIE High Aspect Ratio DOE slides.pptx
HAR Bosch process papers:
1Ultra_Deep_Reactive_Ion_Etching_of_High_Aspect-Ratio_and_Thick_Silicon_Using_a_Ramped-Parameter_Process2018.pdf
2Very_High_aspect_Ratio_Deep_Reactive_lon_Etching_of_Sub_micrometer_Trenches_in_Silicon_2012.pdf
3improvedHAR_hole_Bosch_2022.pdf
Meeting 2/19 summary:
We have run 2 wafers through the DRIE, and have 10 more wafers with the current pattern. We want to write a new mask and pattern new wafers with a new mask, so we have a two-prong plan:
----------------------------------------------------------------------------------------------------------
Current batch:
Some of our wafers came out of the ICP-F cloudy, so we need to strip one of them in the EKC bath and make sure the cloudiness goes away. If it does, then we do not need to strip the rest of the cloudy wafers. If it is still cloudy, we might have bigger problems (micromasking), might not want to use the cloudy wafers.
Before changing any recipe parameters, we are going to run the 3x loopset recipe we have been running, but push to the wafer limits. The first run, we can run with 300 loops on set 3 (1045 total loops). After that, take Dektak/Nanospec data, and then run just the last loopset (1 loop for sets 1 and 2) for 100 or 200 loops, depending on etch rate. We will continue taking the same wafer out and then running the last loopset on it until we have about a 900 um (0.9 mm!) trench. At that point, sacrifice the wafer and take SEM imaging.
For future SEM imaging, in order to not mess up the features during cleaving, make the diamond scribe mark on the backside of the wafer. This mark should also be approaching the trenches from the 5 um side, not the 500 um side.
For the next 9 wafers, we are going to experiment with photoresist backing. We need to find the specific recipe for thickness and bake time, but the bake time is supposedly pretty short (talk to Jean).
As we finish out this batch, we can use a combination of trial and error and DOE to create and optimize a new recipe, but we will probably keep 3 loop sets.
--------------------------------------------------------------------------------------------------------------------------------------
Future batches:
We probably want to write a new mask with many 3 um or 5 um trenches, spaced apart by ~ 50 um. I was thinking about having wide trenches that are easy to cleave, but let me know if this has drawbacks. Using small trenches will maximize the aspect ratio, and the etches will probably not be deep enough to where backside protection is necessary.
For measuring the width/depth of these trenches, we should use an optical microscope to measure the pitch (distance between trenches) and the trench width, and then use SEM imaging to get depth measurements, which we can scale by the pitch measurements. It is super important that all of these measurements are careful to get the best data.
If we do not end up writing a new mask, future batches will need to have backside protection. While photoresist can work, it is more effective to have SiO2 on the backside. We discussed putting 4 um of SiO2 on both sides in the PECVD2, rather than just the front as we are currently doing.
To prevent resist reticulation going forward (likely the cause of the cloudiness), we should change the ICP-F recipe. While the current etch power/time works well for our mask, the wafers may be overheating. To prevent this, we should increase the pressure of our current recipe to 6 mTorr, and introduce loops so that the wafer can cool between etches. We can experiment with this, but a loop with 60 seconds of etching and 30 seconds of resting would work well. By keeping the wafer on the chuck, we can prevent inconsistencies that may occur with loading the wafer multiple times.
-----------------------------------------------------------------------------------------------------------------------------------------
The data for the ICP-F runs we have been doing can help us in a few ways. We can investigate the He flow rate to see if wafers with cloudiness have higher flow rates, potentially indicating a poor clamp and/or poor cooling. More interestingly, we can look at other data to see if there is a clear point where we went from etching SiO2 to etching Si. We expect to see differences in variables such as reflected power, as C4F8 may go from being an etch gas to depositing a polymer.
I am going to collect the dates/times of our previous ICP-F runs, so that Grace can find the data for each wafer. I will put these in a spreadsheet so they are easy to read/find.
-----------------------------------------------------------------------------------------------------------------------------------------
Finally, we talked about the importance of taking data during the etches. While data is stored and can be looked at retroactively, it is much easier and more intuitive to take data while the etch is running. This allows us to spot trends and draw conclusions about what is going on in the process. If we do not have time to take data during the etches, especially for long DRIE runs, we should set up graphs for certain parameters of interest. This can allow us to still see the data in an intuitive way.