## MM_Deep_3x_-3
Recipe name |  | MM_Deep_3x_-3
| Unit | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Step name |  | Strike 1 | Dep 1 | Etch 1_1 | Etch 1_2 | Dep 2 | Etch 2_1 | Etch 2_2 | Dep 3 | Etch 3_1 | Etch 3_2
Process Time | secs | 1 | 1.2 | 1.2 | 2.5 2.5 | 1.2 1.4 | 1.2 1.5 | 2.5 2.5 | 1.4 1.4 | 1.5 1.5 | 2.5 2.5
Process Pressure | mTorr | 40.0 +/- 0% | 40.0 +/- 50% | 30.0 30.0 +/- 50% | 30.0 30.0 +/- 50% | 40.0 40 +/- 50% | 30.0 30 +/- 50% | 30.0 30.0 +/- 50% | 40.0 40.0 +/- 50% | 30.0 30.0 +/- 50% | 30.0 30.0 +/-50%
APC Setpoint Position | % | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
APC Mode |  | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic
Source power | Watts | 2500.0 +/- 0 | 2500.0 +/- 50% | 2500.0 +/- 50% | 2500.0 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50%
Source MU Tune capaciter |  | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
Source MU load capaciter |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source RF Control Mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Source 2 Power | Watts | 500.0 +/- 0 | 500.0 +/- 50% | 500.0 +/- 50% | 500.0 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50%
Source 2 MU tune capaciter |  | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source 2 MU load capaciter |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source 2 RF Control mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Platen HF Power | Watts | 0.0 +/- 0% | 0.0 +/- 0% | 64.0 70.0 +/- 30% | 24 +/- 40% | 0.0 0.0 +/- 0% | 70.0 80 +/- 30% | 24.0 24 +/- 40% | 0.0 0.0 +/- 0% | 80.0 60.0 +/-30% | 20.0 20.0 +/-40$
Platen HF Tune Capaciter Adjust |  | Present | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic
Platen HF Tune Capacter |  | 30.0 +/- 20% | 20.0 +/- 5% | 20.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 20.0 +/- 5% | 20.0 +/- 5% | 20.0 +/-5% | 20.0 +/-5%
Platen HF load Capaciter |  | 60.0 +/- 20% | 30.0 +/- 5% | 30.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5%
Platen HF Padding Capaciter |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Platen HF control mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Platen HF modulation Enabled |  | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable
Platen HF Modulation Frequence | Hz | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150
Platen HF Modulation Duty Cycle | % | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Hellium Pressure | Torr | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/-20 %
Hellium Flow warning level | sccm | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Hellium Flow fault level | sccm | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25
Coil current | Amps | 10.0 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5%
Loop destination |  | 0 | 0 | 0 | 2 | 0 | 0 | 5 | 0 | 0 | 8
Number of loops |  | 0 | 0 | 0 | 262 | 0 | 0 | 483 | 0 | 0 | 550
Loop Variation Parameter |  | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Gas line Config |  | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow
P1 Argon | sccm | 300 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P2 Oxygen | sccm | 100 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P3 Nitrogen | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P4 C4F8 | sccm | 220.0 +/- 20% | 220.0 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 220.0 220 +/-20% | 1.0 1.0 +/-0% | 1.0 1.0 +/- 0% | 220.0 220 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0%
P5 SF6 | sccm | 1.0 +/- 0% | 1.0 +/- 0% | 400.0 400 +/- 20% | 400 +/- 20% | 1.0 1.0 +/-0% | 400.0 400.0+/- 20% | 400.0 400.0 +/- 20% | 1.0 1.0 +/- 0% | 400.0 400.0 +/- 20% | 400.0 400.0 +/- 20%
S1 Oxygen | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
S2 C4F8 | sccm | 110.0 +/- 20% | 110.0 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 110.0 110 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 110.0 110 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0%
S3 SF6 | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%

## MD_Deep_3x_-3 (proposed recipe)
Recipe name |  | MD_Deep_3x_-3
| Unit | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | Blue = proposed change
Step name |  | Strike 1 | Dep 1 | Etch 1_1 | Etch 1_2 | Dep 2 | Etch 2_1 | Etch 2_2 | Dep 3 | Etch 3_1 | Etch 3_2
Process Time | secs | 1 | 1.2 | 1.2 | 2.5 2.5 | 1.2 1.4 | 1.2 1.5 | 2.5 2.5 | 1.4 1.4 | 1.5 1.5 | 2.5 3.5 | Ramping from 2.5 to 3 seconds improved aspect ratio in DOE. Pushing this to 3.5 may improve aspect ratio further.
Process Pressure | mTorr | 40.0 +/- 0% | 40.0 +/- 50% | 30.0 30.0 +/- 50% | 30.0 30.0 +/- 50% | 40.0 40 +/- 50% | 30.0 30 +/- 50% | 30.0 30.0 +/- 50% | 50.0 50.0 +/- 50% | 20.0 20.0 +/- 50% | 20.0 20.0 +/-50% | Research papers + intuition suggest higher passivation pressure & lower etch pressure will lead to higher aspect ratio
APC Setpoint Position | % | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
APC Mode |  | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic
Source power | Watts | 2500.0 +/- 0 | 2500.0 +/- 50% | 2500.0 +/- 50% | 2500.0 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50% | 2500.0 2500 +/- 50%
Source MU Tune capaciter |  | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 4
Source MU load capaciter |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source RF Control Mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Source 2 Power | Watts | 500.0 +/- 0 | 500.0 +/- 50% | 500.0 +/- 50% | 500.0 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50% | 500.0 500 +/- 50%
Source 2 MU tune capaciter |  | 1 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source 2 MU load capaciter |  | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2
Source 2 RF Control mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Platen HF Power | Watts | 0.0 +/- 0% | 0.0 +/- 0% | 64.0 70.0 +/- 30% | 24 +/- 40% | 0.0 0.0 +/- 0% | 70.0 80 +/- 30% | 24.0 24 +/- 40% | 0.0 0.0 +/- 0% | 80.0 60.0 +/-30% | 20.0 20.0 +/-40$ | May want to experiment with platen power, DOE results were unclear (tried ramping 80 --> 60 vs. 80 --> 100)
Platen HF Tune Capaciter Adjust |  | Present | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic | Automatic
Platen HF Tune Capacter |  | 30.0 +/- 20% | 20.0 +/- 5% | 20.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 20.0 +/- 5% | 20.0 +/- 5% | 20.0 +/-5% | 20.0 +/-5%
Platen HF load Capaciter |  | 60.0 +/- 20% | 30.0 +/- 5% | 30.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 25.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5% | 30.0 +/- 5%
Platen HF Padding Capaciter |  | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Platen HF control mode |  | Load | Load | Load | Load | Load | Load | Load | Load | Load | Load
Platen HF modulation Enabled |  | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable | Enable
Platen HF Modulation Frequence | Hz | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150 | 150
Platen HF Modulation Duty Cycle | % | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20 | 20
Hellium Pressure | Torr | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/- 20% | 10.0 10.0 +/-20 %
Hellium Flow warning level | sccm | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10 | 10
Hellium Flow fault level | sccm | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25 | 25
Coil current | Amps | 10.0 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5%
Loop destination |  | 0 | 0 | 0 | 2 | 0 | 0 | 5 | 0 | 0 | 8
Number of loops |  | 0 | 0 | 0 | 262 | 0 | 0 | 483 | 0 | 0 | 550
Loop Variation Parameter |  | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Gas line Config |  | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow | Flow
P1 Argon | sccm | 300 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P2 Oxygen | sccm | 100 +/- 20% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P3 Nitrogen | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
P4 C4F8 | sccm | 220.0 +/- 20% | 220.0 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 220.0 220 +/-20% | 1.0 1.0 +/-0% | 1.0 1.0 +/- 0% | 220.0 220 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0%
P5 SF6 | sccm | 1.0 +/- 0% | 1.0 +/- 0% | 400.0 400 +/- 20% | 400 +/- 20% | 1.0 1.0 +/-0% | 400.0 400.0+/- 20% | 400.0 400.0 +/- 20% | 1.0 1.0 +/- 0% | 400.0 400.0 +/- 20% | 400.0 400.0 +/- 20%
S1 Oxygen | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%
S2 C4F8 | sccm | 110.0 +/- 20% | 110.0 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 110.0 110 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0% | 110.0 110 +/- 20% | 1.0 1.0 +/- 0% | 1.0 1.0 +/- 0%
S3 SF6 | sccm | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5% | 0.0 0.0 +/- 5%